ARISE

Augmented Rehabilitation & Intelligent System for Enhancement

Compliance Dossier

Phase 1, TRL6 demonstration. Class IIa under MDR Annex VIII Rule 11

Documents in this dossier

01   Data Management Plan (DMP)

Engineering specification: data architecture, lifecycle, schemas, pipelines, quality, access

02   Data Protection Impact Assessment (DPIA)

GDPR Article 35 risk assessment, lawful bases, mitigations, twelve-risk register

03   MDR Compliance Plan

Confirmed Class IIa under Rule 11. Annex IX conformity-assessment route. Clinical investigation under MDR Articles 62-82 in Phase 1, Notified Body engagement in Phase 2

Project context

| Funding instrument | ERDF |
| --- | --- |
| Project duration | 18 months, June 2026 to December 2027 |
| Work Package | WP1, Task T1.4 (Compliance Management) |
| Prepared by | Innovina S.r.l. |
| Scientific partner | DINOGMI, University of Genoa |
| Clinical partner | Studio Buccarella |
| Confirmed MDR class | Class IIa under Annex VIII Rule 11 |
| Conformity assessment route | Annex IX (Notified Body required) |
| Dossier version | 2.0, 9 June 2026 |

Contents

01    Data Management Plan

02    Data Protection Impact Assessment

03    MDR Compliance Plan

The three documents in this dossier cross-reference each other. Section numbers refer to the document within which the reference appears. Where one document defers detail to another, the referenced document is named explicitly.

01

Data Management Plan

Engineering specification for ARISE data

Document summary

•  Four storage tiers: edge, hot, warm, cold.

•  Three lifecycle blocks: Phase A (M6 to M15, WP3), Bridge (M8 to M18, WP4), Phase B (M19 to M24, WP5).

•  Five top-level datasets: D-CORPUS (training corpus), D-TRIAL (trial operational), D-USABILITY, D-TELEMETRY, D-ENG.

•  Schemas, pipelines split by phase, quality gates, access control roles.

•  Compliance traceability matrix mapping engineering controls to GDPR and MDR.

## Executive summary

This document is the engineering specification for how ARISE data is captured, processed, stored, accessed, versioned, monitored, retained, and deleted across the project lifecycle.

The plan organises around three layers. The Lifecycle layer describes what happens when, across a pre-trial Phase A in which the training corpus is built and the model is trained, a trial Phase B in which the locked model is used in clinical demonstration, and continuous activities that run for the full project. The Data layer describes the five authoritative datasets, their owners, schemas, and quality gates. The Platform layer describes storage, access, and identity, with technology choices kept generic where the vendor is not yet selected.

GDPR risk assessment lives in the DPIA. MDR-aligned technical documentation lives in the MDR Compliance Plan. This document handles the engineering side of the project.

# Part I: Foundations

## 1. Purpose and scope

### 1.1 Purpose

To specify the engineering controls and operational disciplines that keep ARISE data trustworthy, reproducible, and compliant across the project lifecycle.

### 1.2 Scope

| In scope | Out of scope |
| --- | --- |
| Data inventory, classification, schemas, lineage | GDPR risk assessment (DPIA) |
| Storage, access, identity, encryption | Final cloud vendor and region selection (D2.1B) |
| Pipelines, quality gates, versioning | MDR technical-file structure (MDR Compliance Plan) |
| Ownership and operational responsibilities | UI/UX of the Clinical Dashboard (D2.2) |

### 1.3 Audience

The document is written for the data engineers and platform owners at Innovina who operate the system day to day. It is also intended for the auditors of the funding body, for the Ethics Committee, and for the technical leads at DINOGMI and Studio Buccarella who need to verify that the engineering controls match the partnership agreements.

### 1.4 Technology choices in this document

Storage classes (object store, relational store, cold archive), monitoring service, identity provider, key management service, model registry, and data version-control system are referred to by their generic function rather than by vendor. The specific vendor and region are confirmed in D2.1B.

## 2. Architecture overview

This section describes where data lives and how the system is laid out. The time dimension is covered in Section 3. Detailed schemas are in Section 5 and operational policies are in Sections 6 and 7.

### 2.1 The four storage tiers

| Tier | Function | Latency | Typical contents |
| --- | --- | --- | --- |
| Edge | Real-time inference at the Coach device | Less than 100 ms | On-device inference cache, encrypted local buffer drained to cloud after acknowledgement |
| Hot (operational relational store) | Per-session reads, dashboard queries | Less than 1 s | Sessions, reps, baselines, AE log, annotations |
| Warm (cloud object store) | Per-session JSON, landmark streams, dataset snapshots | Seconds | Corpus snapshots (immutable, versioned), per-session bundles, AE evidence packs |
| Cold (cold-archive object store) | Long-term retention | Hours | Archived corpus, archived trial data, archived AE records |

### 2.2 Where compute runs

| Process | Where | Rationale |
| --- | --- | --- |
| Pose estimation, KPI extraction, classification (live) | Edge (Coach device) | Real-time biofeedback requires sub-100 ms latency and cannot tolerate a cloud round-trip |
| Per-session aggregation, dashboard queries | Cloud (hot tier) | Therapist access between sessions |
| Annotation workflows | Cloud (warm) plus annotator workstations via remote desktop | Annotators never receive a local copy of raw data |
| AI training and evaluation | Cloud (warm) with frozen-snapshot inputs | Reproducibility requires immutable inputs |
| Long-term retention | Cloud (cold) | Cost vs access tradeoff |

## 3. Project lifecycle

ARISE has two distinct data phases bridged by a prototyping period, plus a set of continuous activities that run for the full project duration.

### 3.1 The two phases and the bridge

| Phase | Period | Work package | Goal | Output deliverables |
| --- | --- | --- | --- | --- |
| Phase A. Build the model | M6 to M15 | WP3 (Core AI R&D) | Choose an AI approach, acquire a labelled training corpus, train and evaluate the models | D3.1 (M9, sensor study), D3.2 (M9, AI model research and comparison), D3.3 (M12, hardware design), D3.4 (M15, preliminary AI tests on trained models) |
| Bridge. Build the platform | M8 to M18 | WP4 (Prototyping) | Build the TRL6 prototype (5 to 10 Coach units plus cloud platform) and integrate the locked model into it | D4.1 (M18, TRL6 prototype) |
| Phase B. Validate in clinical use | M19 to M24 | WP5 (TRL6 Validation) | Demonstrate the integrated system on real patients in the relevant clinical environment | D5.1 (M24, Final TRL6 Validation Report) |

The phase boundary is hard. The model is locked at M15 (D3.4). No retraining, no model edit, no taxonomy change happens during the WP4 bridge or WP5 trial. The model that ships in the D4.1 prototype is the same model that is demonstrated in WP5. Any model modification after M15 requires explicit Ethics Committee notification and a re-validation cycle. This rule exists because the TRL6 demonstration validates a specific, identifiable system. Changing the system mid-bridge or mid-trial invalidates the evidence.

A small, distinct annotation activity does happen during Phase B. A sampled subset of trial reps is spot-annotated by DINOGMI to validate the locked model's accuracy against expert ground truth (Task T5.4). This is validation, not training.

### 3.2 Phase A. Build the model (M6 to M15)

Phase A is the WP3 industrial research phase. It begins with sensor selection and AI approach selection, then moves into corpus acquisition, training, and evaluation.

| # | Stage | Period | Producer | Action | Output artefact |
| --- | --- | --- | --- | --- | --- |
| A0 | Sensor selection | M6 to M9 | Innovina + Invenio | Comparative study (RGB, RGB-D, IMU prototypes) | D3.1 Sensor Comparison Study Report |
| A1 | AI approach selection | M6 to M9 | Innovina ML + Invenio | Comparative research of candidate architectures | D3.2 AI model research and comparison |
| A2 | Corpus acquisition | M9 to M15 | Patients and healthy volunteers + clinical sites | RGB capture using the chosen sensor | Raw acquisitions in warm tier (D-CORPUS) |
| A3 | Curation | M9 to M15 | ML lead | Build dataset cut from raw acquisitions | Versioned, immutable snapshot |
| A4 | Annotation, primary | M9 to M15 | DINOGMI annotators | Double-blind labelling of every rep | Annotation records (D-CORPUS, annotations component) |
| A5 | Adjudication | M9 to M15 | DINOGMI senior annotator | Resolve inter-annotator disagreements | Adjudicated labels (D-CORPUS, annotations component) |
| A6 | Quality manifest | M14 to M15 | ML lead | Inter-rater agreement, completeness, per-subject distribution audit | Quality manifest signed off |
| A7 | Model training and evaluation | M12 to M15 | ML pipeline | Train on locked corpus plus adjudicated labels, evaluate on a held-out cut | Trained model artefact plus evaluation metrics |
| A8 | Model lock | M15 | ML lead + scientific lead | Sign-off, register model as locked for TRL6 | D3.4 Preliminary AI Tests, model registered as locked |

In parallel with A0 to A8, T3.5 (Hardware design) runs M9 to M12 and produces D3.3 (Hardware Design Document) at M12, feeding directly into WP4. After A8 the model is immutable. No further training occurs in Phase 1.

### 3.3 Bridge. Build the platform (M8 to M18)

The Bridge is the WP4 prototyping phase. No new clinical data is acquired and no model retraining happens here. The activity is engineering. The team integrates the locked model (D3.4) into the Coach device firmware, builds the Cloud Platform (per D2.1B), and performs end-to-end laboratory tests on the integrated stack.

| # | Stage | Period | Producer | Action | Output artefact |
| --- | --- | --- | --- | --- | --- |
| BR1 | Coach hardware prototyping | M8 to M18 | Innovina Hardware | Build 5 to 10 Coach units from D3.3 design | Hardware units (D4.1 component) |
| BR2 | On-device software | M8 to M18 | Innovina ML + Invenio | Integrate locked model into device firmware, biofeedback layer | Coach firmware build |
| BR3 | Cloud platform build | M13 to M18 | Innovina Platform + Invenio | Build the ingestion API, hot and warm stores, Clinical Dashboard | Cloud platform per D2.1B |
| BR4 | System integration and testing | M16 to M18 | Innovina + Invenio | Laboratory tests on the integrated hardware, software, AI stack | Integration test results |

The artifact at the end of the Bridge is D4.1 (TRL6 prototype). It is the system demonstrated in Phase B.

### 3.4 Phase B. Validate in clinical use (M19 to M24)

| # | Stage | Producer | Action | Output artefact |
| --- | --- | --- | --- | --- |
| B1 | Capture | Patient + Coach device | STS performed in routine therapy session, on-device inference uses the locked model from A8 | Local encrypted buffer (edge) |
| B2 | Ingestion | Coach to ingestion API | Signed TLS upload, schema validation at the API boundary | Rows in hot tier + objects in warm tier |
| B3 | Per-session aggregation | Cloud worker | End-of-session aggregation (KPIs, error counts, quality score) | Session record + per-session bundle |
| B4 | Therapist consumption | Clinical Dashboard | Read by the patient's named therapist | Read-only access |
| B5 | Spot-annotation (T5.4) | DINOGMI annotators on a sampled subset | Compare ARISE output to expert label | Validation records (separate from corpus annotations) |
| B6 | Adverse event tracking | Clinical lead | Log incidents in real time, escalate as required | AE record (D-TRIAL, AE sub-component) |
| B7 | Archive | Lifecycle scheduler | Migrate to cold tier per retention schedule | Cold-tier objects + audit log |

### 3.5 Continuous activities (M1 to M24)

These run uninterrupted across both phases and are not specific to either.

| Activity | Description |
| --- | --- |
| Telemetry | Device health, ingestion rates, latency distributions |
| Logging | Structured application logs, audit logs |
| Monitoring and alerting | Coverage of platform health, paging on threshold violations |
| Backup and snapshots | Nightly backups, point-in-time recovery |
| Retention enforcement | Nightly retention-policy evaluation, deletion jobs |
| Compliance audit | Quarterly reviews of access logs, retention compliance |
| Post-market surveillance preparation | From M18, scaffolding for Article 86 PSUR data collection in Phase 2, given the confirmed Class IIa classification |

# Part II: The data itself

## 4. Data inventory

This section identifies every dataset in scope, who owns it, where the authoritative copy lives, and the consumers it serves.

### 4.1 Authoritative datasets

There are five datasets at the top level. The two most important are D-CORPUS and D-TRIAL. D-CORPUS has two sub-components (raw acquisitions and annotations) because they are logically one dataset consumed together but written by different processes. D-TRIAL has a sub-component (AE log) with stricter access and longer retention than the rest of trial data, but operationally it is the same data stream.

| ID | Dataset | Components | Owner | System of record | Downstream consumers |
| --- | --- | --- | --- | --- | --- |
| D-CORPUS | Training corpus (Phase A) | (a) Raw acquisitions: video, landmarks, KPIs. (b) Annotations: labels, severities, adjudication trail | Innovina ML + DINOGMI | Warm object store (raw, snapshots) + hot relational store (annotations) | Training, evaluation |
| D-TRIAL | TRL6 trial operational data (Phase B) | (a) Per-rep KPIs and sessions. (b) Adverse event log (encrypted with separate key, longer retention) | Innovina Platform. AE sub-component owned by Studio Buccarella | Hot relational store + warm object store | Dashboard, T5.4 analysis, D5.1, DPO and regulator for AE |
| D-USABILITY | Usability questionnaires and interview notes (Phase B) | SUS scores + structured interview transcripts | Innovina Clinical Coordinator + DINOGMI + Studio Buccarella (jointly) | Warm object store (transcripts) + hot relational store (scored) | D5.1 |
| D-TELEMETRY | Device telemetry, system logs | Continuous | Innovina Platform | Monitoring service | On-call ops, post-incident analysis |
| D-ENG | Code, model weights, configs | Continuous | Innovina Engineering | Git (code), model registry (models), data version-control (data pointers) | Development, training |

### 4.2 Classification matrix

Granularity is at the component level where the classification differs across components of the same dataset.

| Dataset, component | Sensitivity | Personal data | Special category |
| --- | --- | --- | --- |
| D-CORPUS, raw acquisitions (video) | Restricted | Yes (biometric) | Yes (health) |
| D-CORPUS, raw acquisitions (landmarks + KPIs) | Confidential | Pseudonymised | Yes |
| D-CORPUS, annotations | Confidential | Linked to subject ID | Yes |
| D-TRIAL, per-rep KPIs | Confidential | Pseudonymised | Yes |
| D-TRIAL, dashboard aggregates | Internal | Pseudonymised | Yes |
| D-TRIAL, AE log | Restricted | Pseudonymised | Yes |
| D-USABILITY | Internal | Pseudonymised | Yes (opinion + health) |
| D-TELEMETRY | Internal | None | No |
| D-ENG | Internal | None | No |

## 5. Schemas

Every dataset that crosses a system boundary has a versioned schema. Schemas are validated at write time (ingestion) and at read time (consumer side) so producers and consumers cannot drift silently.

### 5.1 Per-rep KPI record (D-TRIAL, D-CORPUS)

| Field | Type | Notes |
| --- | --- | --- |
| rep_id | UUID | Stable identifier across all downstream artefacts |
| session_id | UUID | Foreign key to sessions |
| subject_id | string, opaque | Pseudonymised at the point of acquisition |
| device_id | string | Hardware unit identifier |
| acquired_at | timestamp (UTC) | RFC 3339 |
| phase_durations_s | object | One float per phase. Sitting, forward lean, lift-off, standing, stabilisation, descent |
| kpis | object | 18 named numeric KPIs (see D1.1 Section 3) |
| classification | object | { binary, errors[], severities[], confidence } |
| safety_flags | list | Subset of { G1, G2 } if fired |
| pipeline_metadata | object | Inference latency, frame count, dropped frames |

### 5.2 Session record (D-TRIAL)

| Field | Type |
| --- | --- |
| session_id | UUID |
| subject_id | string, opaque |
| started_at, ended_at | timestamps |
| device_id | string |
| therapist_id | string, opaque |
| rep_count, correct_count | integers |
| avg_quality | float |

### 5.3 Annotation record (D-CORPUS, annotations component)

| Field | Type | Notes |
| --- | --- | --- |
| annotation_id | UUID |  |
| rep_id | UUID | Foreign key |
| annotator_id | string, opaque |  |
| label_binary | enum | correct, incorrect |
| label_errors | list | Subset of E1..G2 |
| label_severities | list | Aligned with label_errors, enum { low, medium, high } |
| confidence | enum | low, medium, high |
| free_text_note | string, optional |  |
| created_at | timestamp |  |
| adjudicated | boolean | True if this is the final post-adjudication label |
| adjudicator_id | string, opaque, nullable |  |

### 5.4 Adverse event record (D-TRIAL, AE sub-component)

| Field | Type | Notes |
| --- | --- | --- |
| ae_id | UUID |  |
| subject_id | string, opaque |  |
| session_id | UUID, nullable |  |
| occurred_at | timestamp |  |
| reporter_id | string, opaque |  |
| attributable_to_system | enum | yes, no, unclear |
| severity | enum | minor, serious, critical |
| description | string | Encrypted at rest with a separate key |
| follow_up | text, nullable |  |
| status | enum | open, under_review, closed |

# Part III: Operations

## 6. Pipelines

Pipelines are grouped by the phase in which they run. Phase A pipelines are one-shot and complete before the trial begins. Phase B pipelines run continuously during the trial. Continuous pipelines run across the entire project.

Every pipeline is idempotent and replayable from its inputs. Inputs are addressed by immutable dataset versions. Outputs carry the pipeline version that produced them.

### 6.1 Phase A pipelines (pre-trial, M6 to M15)

| Pipeline | Trigger | Inputs | Outputs | Owner |
| --- | --- | --- | --- | --- |
| Corpus acquisition | Scheduled clinical session | Coach upload | Raw acquisitions in warm tier | Platform + Clinical |
| Annotation manifest build | Manual on dataset cut | Warm slice | Versioned manifest | ML + DINOGMI |
| Annotation labelling | Manifest published | Manifest + clips | Annotation records (D-CORPUS) | DINOGMI annotators |
| Adjudication | Inter-annotator disagreement | Two annotation rows | Adjudicated label (D-CORPUS) | DINOGMI senior |
| Quality manifest | End of annotation | All annotations + clips | Quality manifest (κ scores, completeness, per-subject distribution) | ML lead + DINOGMI |
| Training dataset cut | One-time, end of T3.2 | D-CORPUS (raw + annotations) | Frozen tar with manifest + SHA-256 | ML lead |
| Model training | Frozen dataset + config | Frozen dataset | Model run, model artefact, eval report | ML lead |
| Held-out evaluation | New model artefact | Model + held-out cut | Metrics report | ML lead |
| Model lock and registration | Eval acceptance at M15 | Model artefact + metrics | Registered model marked locked for TRL6 (D3.4) | ML lead + Scientific lead |

### 6.2 Phase B pipelines (trial, M19 to M24)

| Pipeline | Trigger | Inputs | Outputs | Owner |
| --- | --- | --- | --- | --- |
| Trial ingestion | Coach upload during routine session | Raw rep payload | Rows in hot tier + objects in warm tier | Platform |
| Per-session aggregation | End of session | Reps for the session | Session row + per-session bundle | Platform |
| Spot annotation (T5.4) | Validation subset sampled | Subset of trial reps | Validation labels (separate from D-CORPUS annotations) | DINOGMI |
| Accuracy comparison | Spot-annotated subset available | Validation labels + locked-model output | Accuracy report | ML + DINOGMI |
| AE escalation | New AE record | Write to D-TRIAL AE sub-component | Page to clinical lead + DPO, device freeze if attributable | Platform + Clinical |
| Dashboard read paths | Therapist login | Hot tier | Read-only views | Platform |
| Usability data collection | End of session or end of enrolment | Questionnaire submissions | D-USABILITY records | Clinical Coordinator |

### 6.3 Continuous pipelines (M1 to M24)

| Pipeline | Trigger | Function |
| --- | --- | --- |
| Telemetry collection | Continuous | Device health, latency, throughput |
| Log aggregation | Continuous | Structured app logs into searchable index |
| Audit log write | On every privileged access or deletion | Append to immutable audit log |
| Retention enforcement | Nightly | Evaluate retention rules, queue deletions |
| Backup | Nightly | Database backups + object snapshots |
| Quarterly retention audit | Quarterly | Verify nothing past retention remains |

### 6.4 Acquisition flow detail

The acquisition flow is identical in Phase A and Phase B, but consumers differ.

| Step | Component | Output |
| --- | --- | --- |
| Pseudonymisation at enrolment | Studio Buccarella admin | subject_id assigned, mapping held only at the clinical site |
| Per-frame capture | Coach device, edge | RGB frames, 33 landmarks, per-frame features |
| Per-rep aggregation | Coach device, edge | One rep record per repetition |
| Local buffering | Coach device, edge | Encrypted FIFO buffer with disk-quota cap |
| Ingestion | Cloud API | Schema validation, signed write |
| Acknowledgement | Cloud API to Coach | After durable write. Device purges local copy only after ACK |

## 7. Data quality

### 7.1 Dimensions tracked

| Dimension | Definition | Measurement |
| --- | --- | --- |
| Completeness | Fraction of expected fields present | Per-record schema validation, per-batch metric |
| Accuracy (vs annotation) | Agreement of automated classification with annotated ground truth | Held-out test set in Phase A, spot-annotated subset in Phase B |
| Consistency | FK integrity, phase boundaries match rep boundaries | Nightly consistency job |
| Timeliness | Time from acquisition to availability in warm tier | Per-event metric |
| Validity | Values within expected ranges (angles in [0, 180], KHR in [0, 5], etc.) | Schema constraints + range checks |
| Uniqueness | rep_id unique across the corpus | Database constraint |

### 7.2 Quality gates

A dataset cannot be promoted across a boundary unless it passes its gate.

| Promotion | Gate |
| --- | --- |
| Edge to cloud | Schema validation, signed write, ACK |
| Cloud to annotation (Phase A) | Manifest contains all expected reps, no NaN KPIs above tolerance |
| Annotation to adjudicated | Both annotators have labelled, disagreement routed to adjudicator |
| Adjudicated to training cut | Inter-rater agreement target (Cohen κ greater than or equal to 0.70 binary, Fleiss κ greater than or equal to 0.60 per code) on calibration subset |
| Training cut to training run | SHA-256 of manifest matches expected, size and label distribution within tolerance |
| Trained model to lock | Held-out F1, error-category macro F1, per-stratum bias metrics meet acceptance |

### 7.3 Anomaly detection and triage

| Signal | Action |
| --- | --- |
| Spike in schema validation failures | Page on-call, freeze ingestion until root cause known |
| Sudden drop in average inference confidence | Auto-open ticket, sample affected reps to QA queue |
| Drift in KPI distributions (per device, site, cohort) | Weekly drift report, manual triage |
| Per-subject label outlier distribution (Phase A) | Auto-flag for re-review (the "audit before calibrating" rule) |
| AE attributable to system (Phase B) | Immediate page to DPO + clinical lead, freeze affected device |

# Part IV: Platform

## 8. Storage, access, identity

### 8.1 Storage layout

The cloud has two object-store tiers (warm and cold) and one relational store (hot). Within each tier, data is organised by purpose. The vendor-specific path syntax is not in scope here because the vendor is not yet selected.

Warm tier (cloud object store). Holds operational and corpus artefacts. Each top-level area contains a subtree by purpose. The training corpus area is immutable and version-locked, with one folder per dataset version. The trial area is organised by site and acquisition date, with per-session JSON files and per-rep landmark files. The adverse-event area is organised by date and uses a separate encryption key, so even a platform engineer with general warm-tier access cannot read AE descriptions without the AE key. The usability area is organised by date and subject identifier.

Cold tier (cold-archive object store). Holds long-term archives. The corpus archive mirrors the warm-tier dataset-version structure for retained snapshots. The trial archive is partitioned by year. The AE archive is partitioned by year. A separate area, "raw video pending delete," holds raw video that has not yet been deleted and is auto-purged on annotation sign-off.

Hot tier (relational store). Holds five schemas. The "core" schema holds sessions, reps, and baselines. The "annotations" schema holds the corpus annotation rows. The "ae" schema holds adverse events. The "usability" schema holds the scored questionnaires. The "audit" schema holds the access and deletion logs.

### 8.2 Roles

The access model uses a small, declarative role set held in version control.

| Role | Capabilities |
| --- | --- |
| therapist | Read patient data within own caseload, create AE entries |
| annotator | Read corpus clips, write annotation rows |
| adjudicator | Read annotation rows, write adjudicated rows |
| engineer | Read versioned dataset snapshots, trigger training jobs, write to model registry, operate infrastructure, trigger deletion jobs. Cannot read the encrypted AE description field |
| dpo | Read audit log, trigger erasure for a subject |
| security_auditor | Read audit log only |

Privileged roles require MFA. Privileged sessions are time-boxed to a maximum of four hours.

### 8.3 Access patterns

| Persona | Path | Mechanism |
| --- | --- | --- |
| Therapist | Clinical Dashboard | Identity provider with MFA, backend enforces row-level filter to patient cohort |
| Annotator | Annotation workstation | Remote desktop into bastion, no local data persistence, MFA + just-in-time access |
| Engineer | Notebook environment or operational tools | SSO with MFA, scoped to versioned dataset snapshots, break-glass procedure logged separately |
| DPO | Audit tools | SSO with MFA, read-only audit log |
| External auditor | Read-only audit slice | Time-boxed credentials, full access logging, signed agreement |
| Notified Body auditor | Read-only audit slice + QMS documentation | Time-boxed credentials, signed confidentiality agreement, scope limited to the conformity assessment of the Class IIa device per MDR Annex IX |

## Part V: Compliance

## 9. Compliance traceability matrix

Each engineering control in this document supports one or more regulatory obligations. This table is the audit-traceability artefact.

| Control | GDPR | MDR |
| --- | --- | --- |
| Pseudonymisation at acquisition (Section 6.4) | Art. 32(1)(a) | Annex II Section 6, Annex IX Section 2.2 (traceability for Class IIa) |
| Encryption at rest and in transit (Section 8) | Art. 32(1)(a) | Annex I GSPR 17.2 cybersecurity, Annex II Section 6 |
| Audit log of access and deletions (Section 6.3) | Art. 30, Art. 32(1)(d) | Annex II Section 6, traceability evidence for Annex IX audit |
| Retention schedule and deletion | Art. 5(1)(e), Art. 17 | Article 10(8) ten-year retention for Class IIa technical documentation, Annex IX Section 8 |
| Quality gates (Section 7.2) | Art. 5(1)(d) accuracy | Annex II Section 6.2 verification and validation |
| Subject erasure procedure | Art. 17, Art. 28(3)(g) | n/a |
| Bias monitoring across strata (Section 7.3) | Recital 71 (automated decisions) | Annex I GSPR 17 (software), Annex I GSPR 22 (devices intended for use by lay persons not applicable here) |
| KPI-distribution drift monitoring (Section 7.3) | n/a | Article 88 trend reporting in Phase 2 post-market |

GDPR risk weighting and mitigation depth are in the DPIA. MDR-specific record-keeping requirements are in the MDR Compliance Plan. This document is the engineering substrate that both rely on.

02

Data Protection Impact Assessment

GDPR Article 35 mandatory assessment

Document summary

•  DPIA mandate under GDPR Article 35(3) trigger (a) automated decision-making, (b) special category at scale, (c) novel technology.

•  Lawful bases summary across Article 6 and Article 9, with explicit non-reliance on 6(1)(d to f) and 9(2)(g, i).

•  Twelve identified risks scored on likelihood and severity, mapped to project phases.

•  Inline operational controls: retention schedule, deletion procedure, audit logging, breach response, encryption.

•  Article 80 of MDR applies in Phase 1 (clinical investigation), Articles 83 to 87 take over in Phase 2 for the confirmed Class IIa device.

## Executive summary

This document is the Data Protection Impact Assessment for the ARISE project, prepared under GDPR Article 35. A DPIA is mandatory for ARISE on three independent grounds. The processing involves a systematic and extensive evaluation of individuals based on automated processing (AI inference). It involves special category data (health) on a large scale. It involves new technologies that have not previously undergone DPIA in the same context.

The assessment is organised around the project lifecycle defined in the Data Management Plan. There is a pre-trial Phase A (M6 to M15, WP3) in which the training corpus is built and the model is trained, a Bridge period (M8 to M18, WP4) in which the locked model is integrated into the prototype platform, and a Phase B (M19 to M24, WP5) in which the system is demonstrated on real patients. Risk profiles differ across phases and are mapped accordingly.

Twelve risks have been identified, with pre-mitigation likelihood and severity scored on a three-point scale. After the mitigations in Section 6 are in place, all residual risks fall to low likelihood and low to medium severity. No residual risk crosses the threshold for prior consultation with the Garante per la protezione dei dati personali under Article 36.

The DPIA covers GDPR-specific operational controls (retention, deletion, audit logging, encryption, breach response) inline in Section 6. Wider engineering controls (storage, access, pipelines, quality gates) are specified in the DMP and referenced where relevant.

The medical-device classification under MDR has been confirmed as Class IIa under Annex VIII Rule 11. This DPIA reflects that confirmation. The Article 22 safeguard in Section 6.8 is calibrated accordingly: the locked model produces information used by the supervising clinician, who remains the clinical decision-maker, so the processing does not constitute a fully automated decision producing legal or similarly significant effects on the data subject.

# Part I: Foundations

## 1. Purpose and legal basis for the DPIA

### 1.1 Why a DPIA is mandatory for ARISE

GDPR Article 35(1) requires a DPIA when processing is likely to result in a high risk to the rights and freedoms of natural persons. Article 35(3) lists three case-based triggers. ARISE meets all three independently.

| Trigger (Article 35(3)) | How it applies to ARISE |
| --- | --- |
| (a) Systematic and extensive evaluation of personal aspects based on automated processing | The system performs AI-driven classification of repetitions and identifies clinical errors used to inform therapy decisions |
| (b) Processing of special category data (Article 9) on a large scale | Health data is processed across the training corpus and the TRL6 trial, covering hundreds of subjects |
| (c) Use of new technologies | Markerless RGB-based 3D pose estimation combined with biomechanical AI inference is novel in this clinical context |

### 1.2 Scope of this DPIA

This DPIA covers the entire ARISE Phase 1 processing activity from corpus acquisition (M6) through TRL6 clinical demonstration (M24).

| In scope | Out of scope |
| --- | --- |
| Processing during Phase A, Bridge, Phase B | Phase 2 commercialisation processing (separate Phase 2 DPIA at CE marking) |
| Personal-data streams D-CORPUS, D-TRIAL, D-USABILITY (per DMP Section 4.1) | Non-personal streams D-TELEMETRY and D-ENG |
| Risks to data subject rights and freedoms | Risks to clinical or commercial outcomes (covered by the MDR Compliance Plan) |
| Privacy-specific operational controls (retention, deletion, audit, breach) included inline in Section 6 | Wider engineering controls in the DMP |

### 1.3 Audience

The document is intended for the Innovina Data Protection Officer and Project Lead. It is also reviewed by the scientific lead at DINOGMI, by the clinical lead at Studio Buccarella, by the Ethics Committee in connection with the protocol submission, and by ERDF compliance auditors.

## 2. Lawful bases summary

The processing rests on the following GDPR Article 6 and Article 9 bases. Granular explicit consent is the principal basis. Secondary bases support specific operations that retain a regulatory or contractual character beyond the consent envelope.

| Activity | Article 6 basis | Article 9 basis |
| --- | --- | --- |
| Corpus acquisition (Phase A) | 6(1)(a) explicit consent | 9(2)(a) explicit consent, with 9(2)(j) scientific research (Art. 89(1) safeguards) as secondary basis |
| Model training (Phase A) | 6(1)(a) explicit consent (same data subjects, same consent) | 9(2)(a) and 9(2)(j) |
| Trial operation, therapeutic context (Phase B) | 6(1)(a) explicit consent + 6(1)(b) contract for care between patient and clinic | 9(2)(a) and 9(2)(h) provision of health care |
| Adverse event log handling (Phase B) | 6(1)(c) legal obligation under MDR Article 80 (recording and reporting of adverse events during clinical investigations) | 9(2)(a) explicit consent (the participant consents to AE recording at enrolment), with 9(2)(h) as secondary basis (care-relationship context) |
| Usability questionnaires (Phase B) | 6(1)(a) explicit consent | 9(2)(a) |
| Sharing with DINOGMI for research | 6(1)(a) | 9(2)(a), 9(2)(j) |
| Sharing with Studio Buccarella for care | 6(1)(b) | 9(2)(h) |

Article 6(1) bases used are (a) consent, (b) contract, and (c) legal obligation. Article 9(2) bases used are (a) explicit consent, (h) provision of health care, and (j) scientific research with the safeguards required by Article 89(1).

The MDR reference for Phase 1 is Article 80 (clinical investigation AE recording), not Article 87 (post-market vigilance), because the device is not yet placed on the Union market during the TRL6 demonstration. Article 87 becomes the controlling reference in Phase 2 after CE marking, with the post-market surveillance framework of Articles 83 to 86 applying to the Class IIa device.

Consent is documented in writing through the Informed Consent procedure that is attached to D1.2. Consent is granular: the subject consents separately to therapeutic use, to use of data for AI training, and to use for scientific publication and dissemination. Withdrawal of consent is operationalised through the deletion procedure in Section 6.5.

# Part II: The processing

## 3. Description of the processing

### 3.1 Nature

The processing comprises four activities that vary by phase. Engineering details are in the DMP. This section restates the privacy-relevant aspects only.

| Activity | Phase | Privacy-relevant aspect |
| --- | --- | --- |
| Video capture of patient performing STS | Phase A and Phase B | Biometric and health data, special category (Article 9) |
| Derivation of 3D anatomical landmarks and biomechanical KPIs | Phase A and Phase B | Derived special category data |
| Classification by rule layer and AI model | Phase A (laboratory) and Phase B (in operation, with locked model from D3.4) | Automated processing of special category data |
| Annotation and adjudication of the corpus | Phase A only (M9 to M15) | Human-attributed labels on identified biometric data |

### 3.2 Scope

| Aspect | Detail |
| --- | --- |
| Categories of data subject | Adult patients in active rehabilitation, healthy adult volunteers for the AI training corpus |
| Categories of data | Biometric (3D pose), health (rehabilitation performance), opinion (usability questionnaires) |
| Volume estimate | Approximately 50 to 100 subjects in Phase A, 50 to 150 subjects in Phase B |
| Duration of active processing | M6 to M24 |
| Retention after active processing | See Section 6.4 (retention schedule) |

### 3.3 Context

| Aspect | Detail |
| --- | --- |
| Relationship with data subjects | Patient and clinical care provider (Studio Buccarella), or research participant and research institution (DINOGMI) |
| Reasonable expectation of processing | Yes. Consent is informed, explicit, granular |
| Vulnerable populations | Yes. Geriatric patients, post-stroke patients with possible cognitive impairment. Special safeguards apply (Section 6.6) |
| Previously processed elsewhere | No |
| Public concerns | None specifically identified to date |

### 3.4 Purposes

Each purpose has a necessary justification.

| Purpose | Necessity justification |
| --- | --- |
| Real-time biofeedback to the patient during the rep | Core therapeutic function. Cannot be achieved without per-frame biometric capture |
| Longitudinal therapeutic monitoring by the therapist | Requires per-rep KPI data. Pseudonymisation is sufficient |
| Training AI models that generalise across patient populations | Requires diverse labelled examples. Pseudonymised landmark and label data is sufficient |
| TRL6 validation evidence (D5.1) | Required by the funding programme and by the Ethics Committee |
| Post-market surveillance preparation (MDR alignment for the confirmed Class IIa device) | Required for Phase 2 commercialisation pathway under MDR Articles 83 to 86 |

## 4. Necessity and proportionality

### 4.1 Data minimisation

The dataset retained long-term is the minimum necessary for the stated purposes.

| Minimisation measure | Operational detail |
| --- | --- |
| Raw video deleted after labelling sign-off | Retention 6 to 12 months from acquisition. Trigger is annotation manifest sign-off |
| Identifier mapping held only at the clinical site | Innovina engineers and annotators never receive the mapping |
| AI training on pseudonymised data only | Identifiers removed at the corpus snapshot boundary |
| Per-rep aggregation | No continuous physiological streams retained beyond the rep |

### 4.2 Accuracy

Annotator double-labelling and DINOGMI adjudication ensure label accuracy (DMP Section 3.1 and Section 6.1). Data subjects can request correction of inaccurate data through the channel described in Section 4.4 below.

### 4.3 Information provided to data subjects

Under Articles 13 and 14, each data subject receives a structured Information Sheet at recruitment. The Information Sheet is finalised in D1.2 and covers the following elements.

| Element | Notes |
| --- | --- |
| Identity of the Controller | Named in the Information Sheet |
| Purposes and lawful basis | Granular: therapeutic, training, publication |
| Recipients | Innovina, DINOGMI, Studio Buccarella |
| Retention | Per Section 6.4 below |
| Rights | Access, rectification, erasure, restriction, portability, object, not subject to fully automated decision-making |
| Right to lodge a complaint | With the Garante per la protezione dei dati personali |
| Automated processing safeguard | No fully automated clinical decision is taken on the patient. ARISE is a Class IIa medical device that produces information for the therapist, who remains the clinical decision-maker |

### 4.4 Data subject rights

Rights are operationalised through the clinical site, who relays through the subject ID mapping to Innovina when needed.

| Right (GDPR) | Mechanism |
| --- | --- |
| Information (Art. 13, 14) | Information Sheet at recruitment |
| Access (Art. 15) | Patient requests via Studio Buccarella |
| Rectification (Art. 16) | Same channel |
| Erasure (Art. 17) | Same channel. Technical procedure in Section 6.5 below |
| Restriction (Art. 18) | Same channel |
| Portability (Art. 20) | Limited applicability. Available on request in a machine-readable format |
| Object (Art. 21) | Same channel. Effective from receipt |
| Not subject to automated decisions (Art. 22) | Confirmed. ARISE is decision support. The therapist decides |

### 4.5 Joint controller and processor arrangements

| Agreement | Parties | Scope |
| --- | --- | --- |
| Data Processing Agreement (Art. 28) | Innovina (Processor) and Studio Buccarella (Controller) | Clinical operation during Phase B |
| Joint Controller Agreement (Art. 26) | Innovina and DINOGMI | AI training corpus (Phase A) |

Both agreements include sub-processor obligations, security commitments, audit rights, and breach notification timelines.

# Part III: Risk and mitigation

## 5. Risk register

### 5.1 Identified risks

Twelve risks have been identified across the project lifecycle. Each is scored on likelihood and severity using a three-point scale (Low, Medium, High) before mitigation. The phase in which the risk is most prominent is indicated. Affected rights are expressed in concrete GDPR terms: privacy, self-determination, autonomy, non-discrimination.

| ID | Risk | Phase most exposed | Likelihood (pre) | Severity (pre) | Affected rights |
| --- | --- | --- | --- | --- | --- |
| RD1 | Re-identification of pseudonymised data via biometric pose signature | A, B | Medium | Medium | Privacy, non-discrimination |
| RD2 | Unauthorised access to clinical data at rest | A, B, Bridge | Low | High | Privacy, self-determination |
| RD3 | Unauthorised access in transit | A, B | Low | High | Privacy |
| RD4 | Data breach with public disclosure | A, B, Bridge | Low | High | Privacy, self-determination, non-discrimination |
| RD5 | Function creep: data used for purposes not consented to | A, B | Medium | High | Self-determination, autonomy |
| RD6 | Retention beyond schedule | A, B | Medium | Medium | Self-determination |
| RD7 | Bias in AI training data leading to inequitable care | A | Medium | Medium | Non-discrimination |
| RD8 | Patient inferred clinical condition disclosed beyond intended audience | B | Low | High | Privacy, possible economic harm |
| RD9 | Misclassification by the locked model leading to harmful therapeutic recommendation | B | Low | Medium | Health |
| RD10 | Adverse-event log mishandled, exposing identifiable patient | B | Low | High | Privacy |
| RD11 | Continued processing of a withdrawn participant's data due to incomplete propagation of the withdrawal across systems | A, B | Medium | Medium | Self-determination |
| RD12 | Re-identification of an adverse-event patient via pose signature stored alongside the AE record | B | Low | High | Privacy |

### 5.2 Risk-by-phase summary

| Phase | Risks predominantly active in this phase |
| --- | --- |
| Phase A (corpus, M6 to M15) | RD1, RD5, RD7, RD11 |
| Bridge (prototype, M8 to M18) | RD2, RD4 |
| Phase B (trial, M19 to M24) | RD8, RD9, RD10, RD12, plus continuing RD1 to RD7 |
| Continuous (M1 to M24) | RD2, RD3, RD4, RD6, RD11 |

## 6. Mitigation measures

Each risk in Section 5 is mapped to one or more mitigations below. Wider engineering controls live in the DMP. Privacy-specific operational controls (retention, deletion, audit, breach, encryption) are stated inline here because they are GDPR controls and belong in this document.

### 6.1 Against re-identification (RD1, RD12)

| Measure | Detail |
| --- | --- |
| Identifier mapping held only at the clinical site | Innovina never receives the mapping (DMP Section 6.4) |
| Pose data exported as a 23-joint BVH-compatible skeleton, not raw biometric template | D1.1 Section 2.2 |
| Raw video never shared outside Innovina, deleted at retention trigger | Section 6.4 and Section 6.5 below |
| Annotation guidelines exclude rare-condition outliers that could be re-identified by inference | DMP Section 3.1 |
| AE record stores no raw pose data, only a textual description and pointers to the relevant session and rep | Section 6.9 below |

### 6.2 Against unauthorised access (RD2, RD3): encryption and access control

| Measure | Detail |
| --- | --- |
| Encryption at rest | All personal-data tiers use encryption with customer-managed keys |
| Encryption in transit | Transport Layer Security on every data-bearing connection |
| Role-based access, least privilege | Roles defined in DMP Section 8.2, access patterns in DMP Section 8.3 |
| Time-boxed elevation of privileged roles | Maximum 4 hours per session for engineering roles |
| Audit log of all access events | Append-only, immutable |
| Penetration testing | First test before WP4 backend goes live (M15), annual thereafter |

### 6.3 Against breach (RD4): breach response procedure

| Measure | Detail |
| --- | --- |
| Breach detection | Monitoring alerts on anomalous access patterns |
| Supervisory authority notification | Within 72 hours of confirmed breach (Article 33). DPO is the notifier |
| Data subject notification | Per Article 34 when the breach is likely to result in high risk to rights and freedoms. Channel via the clinical site |
| Documentation | Incident log entry. Root cause analysis |
| Annual simulation | One tabletop exercise per year |

### 6.4 Against retention beyond schedule (RD6): retention schedule

Retention is enforced automatically by a nightly job that evaluates retention rules against the audit log.

| Dataset, component | Retention | Justification |
| --- | --- | --- |
| D-CORPUS, raw acquisitions (video) | 6 to 12 months from acquisition | Time needed to derive and audit the landmark stream |
| D-CORPUS, raw acquisitions (landmarks + KPIs) | 10 years from study end | MDR Article 10(8) minimum retention for the technical documentation of a Class IIa device |
| D-CORPUS, annotations | 10 years | Auditability of training corpus |
| D-TRIAL, operational (per-rep KPIs, sessions) | 10 years from session | MDR Article 10(8) minimum retention for Class IIa |
| D-TRIAL, AE log sub-component | 15 years from incident | Medical-device incident record-keeping |
| D-USABILITY | 10 years | MDR Article 10(8) minimum retention for Class IIa |
| Subject ID mapping (held only at clinical site) | 10 years | Required for data subject rights exercise |

Retention for non-personal data (D-TELEMETRY, D-ENG) is governed by the DMP and not in scope here.

### 6.5 Against function creep (RD5) and withdrawn-data propagation (RD11)

| Measure | Detail |
| --- | --- |
| Granular consent | Separate boxes on the consent form for therapeutic use, AI training, publication. Documented in D1.2 |
| Purpose-specific access control | AI training environment cannot access clinical-only data, and vice versa (DMP Section 8) |
| Erasure procedure | On withdrawal of consent, the DPO triggers the erasure job, which propagates the subject_id to deletion across hot, warm, and cold tiers, and removes the subject from any open dataset cut not yet frozen. Anonymised aggregates already used in publications cannot be retracted (this is disclosed in the Information Sheet) |
| Erasure verification | Quarterly audit confirms that no subject_id flagged for erasure remains in any tier |

### 6.6 Against bias (RD7)

| Measure | Detail |
| --- | --- |
| Stratified recruitment | Across age, gender, condition type. Documented in D1.2 |
| Bias metrics in D5.1 | Reported per demographic stratum |
| Annotation audit | DMP Section 7.3 flags subjects whose labels show systematic outlier distribution before training |

### 6.7 Against inferred-condition disclosure (RD8)

| Measure | Detail |
| --- | --- |
| Pseudonymisation | DMP Section 6.4 |
| Strict role-based access at the Dashboard | Only the patient's named therapist sees their data (DMP Section 8.2) |

### 6.8 Against misclassification harm (RD9)

| Measure | Detail |
| --- | --- |
| Article 22 safeguard | The locked model produces information for the therapist, who remains the clinical decision-maker. No fully automated clinical decision is made on the patient. This is consistent with the confirmed Class IIa classification of the device (decision-support software for therapeutic decisions) |
| Biofeedback limitation | Real-time biofeedback is limited to motor-cue guidance and is not a clinical diagnosis (D1.1) |
| Always-on safety flags | G1 and G2 fire regardless of phase, calibration, or personalisation (D1.1 Section 6.6) |
| Adverse event log | Captures any incident attributable to system use |

### 6.9 Against AE log mishandling (RD10)

| Measure | Detail |
| --- | --- |
| Restricted access | Clinical lead at the trial site and Innovina |
| Pseudonymisation at entry | AE records carry the opaque subject_id, never the patient's real identity |
| Separate-key encryption | AE description field encrypted with a key not held by general engineering roles |
| No raw pose data attached | AE record links by session_id and rep_id to indices, not to raw landmarks (mitigation against RD12) |

### 6.10 Audit logging policy

Two logs are maintained, both append-only and immutable.

| Log | Retention | Scope |
| --- | --- | --- |
| Data-access audit log | 2 years | Every read of personal data, with timestamp, role, subject_id (pseudonymised), action |
| Deletion audit log | 10 years | Every deletion event, with timestamp, role, dataset, justification |

### 6.11 Residual risk summary

After the mitigations in this section are in place, all residual risks fall to low likelihood and low to medium severity. No residual risk crosses the threshold for prior consultation with the supervisory authority under Article 36.

# Part IV: Governance

## 7. Roles and consultations

| Role | Organisation | Responsibility under this DPIA |
| --- | --- | --- |
| Data Protection Officer | Innovina | Owns this DPIA, approves and reviews, primary point of contact for the supervisory authority |
| Project Lead | Innovina | Operational responsibility for implementing mitigations |
| Scientific Lead | DINOGMI | Approves mitigations relevant to research processing (Section 6.6 bias) |
| Clinical Lead | Studio Buccarella | Approves mitigations relevant to clinical processing (Section 6.7 inferred condition, Section 6.9 AE) |
| Joint Controller | Innovina and DINOGMI | Article 26 agreement governs the training corpus |
| Patient representatives | Via Studio Buccarella where available | Consulted during design of the Information Sheet and consent procedure |

## 8. Approval and review cadence

| Event | Trigger | Outcome |
| --- | --- | --- |
| Material change review | Any material change to purpose, data type, recipients, retention, security | New DPIA version |
| Breach review | Any data breach | New DPIA version with updated risk register |
| Routine review | At M12 (project mid-term) and M24 (project end) | Confirmation or revision |
| Phase transition review | Before Phase A starts (M6), before Phase B starts (M19) | Confirmation that controls are in place |
| Pre-Phase 2 review | At the end of Phase 1, before CE marking is sought for the Class IIa device | A Phase 2 DPIA is opened that supersedes the present document for the commercial post-market processing |

03

MDR Compliance Plan

Medical-device regulatory alignment under EU 2017/745. Confirmed Class IIa

Document summary

•  Classification confirmed as Class IIa under MDR Annex VIII Rule 11. The companion document ARISE Classification Scenarios is now archived.

•  Conformity-assessment route: Annex IX (full QMS assessment plus assessment of the technical documentation). Notified Body involvement required.

•  Phase 1 clinical investigation under MDR Articles 62 to 82, AE recording per Article 80.

•  Phase 2 post-market obligations include Article 86 PSUR every two years, Article 87 vigilance, mandatory PMCF per Annex XIV Part B.

•  Full ISO 13485 quality management system and ISO 14971 risk management process scoped for Class IIa.

•  Notified Body engagement plan: shortlist at M14, pre-submission at M18 to M20, contract Q1 Phase 2, QMS audit Q2 Phase 2.

## Executive summary

This document specifies how the ARISE project aligns its development, validation, and documentation practices to Regulation (EU) 2017/745 (MDR), so that the Phase 1 demonstration produces evidence reusable for Phase 2 CE marking without rework.

The medical-device classification of ARISE under MDR Annex VIII has been confirmed as Class IIa under Rule 11. The previously circulated companion document ARISE Classification Scenarios is now superseded by the present plan and is archived. All class-dependent fields in this document are populated accordingly.

Phase 1 does not require CE marking. Its demonstration takes place under an approved Ethics Committee protocol (D1.2) following the clinical-investigation framework of MDR Articles 62 to 82. However, every artefact produced during Phase 1 (technical documentation, risk register, clinical-evaluation evidence, post-market surveillance plan, quality system records) is structured to be MDR-traceable for the Class IIa conformity assessment that follows in Phase 2.

The chosen conformity-assessment route for Phase 2 is MDR Annex IX (full quality management system assessment plus assessment of the technical documentation), on the basis that it provides the broader QMS validation that is useful for the Equipment-as-a-Service commercial model. Notified Body involvement is required for Class IIa under Annex IX. The Notified Body engagement is planned to begin in the second half of Phase 1, with formal QMS audit in the early part of Phase 2.

# Part I: Foundations

## 1. Purpose and scope

### 1.1 Purpose

To ensure that Phase 1 produces MDR-traceable documentation, so that Phase 2 commercialisation (CE marking) is not delayed by retroactive regulatory work, and to specify the MDR obligations that apply to the Phase 1 clinical investigation under the confirmed Class IIa classification.

### 1.2 Scope

| In scope | Out of scope |
| --- | --- |
| Device identification, intended purpose | GDPR risk assessment (DPIA) |
| MDR classification and conformity assessment pathway | Engineering implementation detail (DMP) |
| Risk management process | UI/UX of the Coach and Dashboard (D2.2) |
| Clinical evaluation strategy (full evaluation per Class IIa requirements) | Project management (T1.1) |
| Technical documentation structure (Annex II and Annex III) | Final Notified Body selection (decided in Phase 2 from the agreed shortlist) |
| Vigilance under MDR Article 80 (Phase 1) and Articles 83 to 87 (Phase 2) |  |
| Quality management alignment with ISO 13485 |  |

## 2. Regulatory framework

The applicable regulation is Regulation (EU) 2017/745 (MDR), which is directly applicable across the European Union. ARISE is in clinical investigation during Phase 1 (MDR Articles 62 to 82), not in post-market vigilance. This distinction governs which obligations apply when.

| Lifecycle stage | Project phase | Applicable MDR provisions |
| --- | --- | --- |
| Clinical investigation (no CE marking) | Phase 1 (M1 to M24) | Articles 62 to 82, Annex XV. AE recording under Article 80 |
| Conformity assessment | Phase 2 (post-Phase 1) | Articles 52 to 60, Annex IX (chosen route for Class IIa) |
| Placing on market | Phase 2 (post CE marking) | Articles 5 to 13, Annex I (general safety and performance requirements), Annex III (post-market surveillance plan) |
| Post-market surveillance | Phase 2 ongoing | Articles 83 to 86, with Article 86 PSUR cadence for Class IIa devices |
| Post-market vigilance | Phase 2 ongoing | Articles 87 to 92 |

# Part II: Device and classification

## 3. Device description and intended purpose

### 3.1 Device identification

| Attribute | Value |
| --- | --- |
| Trade name | ARISE Coach (hardware + on-device software) and ARISE Clinical Dashboard (cloud software) |
| Manufacturer | Innovina S.r.l. |
| Configuration | All-in-one camera plus edge computing unit at the clinical site, plus cloud Dashboard accessed via authenticated web client |
| Components | Coach device (camera, edge inference unit, biofeedback display), Cloud Dashboard (secure web platform), locked AI model from D3.4 |
| Software safety class per IEC 62304 | Class B (no death or serious injury is possible from software failure, given mandatory clinician supervision and the contactless nature of the device) |
| Unique Device Identifier (UDI) | Assigned in Phase 2 prior to CE marking, registered in EUDAMED |

### 3.2 Intended purpose statement

ARISE is intended to support rehabilitation of adult patients performing the Sit-to-Stand exercise by providing real-time biofeedback to the patient in the form of visual and auditory cues during the repetition, and by providing objective biomechanical analytics to the supervising therapist in the form of per-rep KPIs, error counts, longitudinal trends, and safety alerts.

Intended user populations are adult rehabilitation patients supervised by licensed physiotherapists. The intended use environment is clinical rehabilitation centres, gyms, and care facilities.

The information produced by ARISE is intended to support therapeutic decisions made by the supervising clinician. The device does not replace clinical judgment.

## 4. MDR classification

### 4.1 Confirmed classification

| Field | Value |
| --- | --- |
| Confirmed class | Class IIa |
| Applicable classification rule | MDR Annex VIII, Rule 11 |
| Conformity assessment route | Annex IX (full QMS assessment plus assessment of the technical documentation) |
| Notified Body involvement required | Yes, mandatory for Class IIa under Annex IX |
| Software safety class per IEC 62304 | Class B |
| Classification authority | Regulatory function of the University of Genoa, in consultation with Innovina S.r.l. as legal manufacturer |
| Date of confirmation | June 2026 |

### 4.2 Rule 11 application and rationale

MDR Annex VIII Rule 11 states:

Software intended to provide information which is used to take decisions with diagnosis or therapeutic purposes is classified as class IIa, except if such decisions have an impact that may cause death or an irreversible deterioration of a person's state of health, in which case it is in class III, or a serious deterioration of a person's state of health or a surgical intervention, in which case it is classified as class IIb. Software intended to monitor physiological processes is classified as class IIa, except if it is intended for monitoring of vital physiological parameters, where the nature of variations of those parameters is such that it could result in immediate danger to the patient, in which case it is classified as class IIb. All other software is classified as class I.

The application to ARISE is as follows.

| Step | Question | Answer |
| --- | --- | --- |
| 1 | Is ARISE software within MDR scope as a medical device under Article 2(1)? | Yes. The intended purpose includes the treatment and alleviation of motor disability through rehabilitation support, which is a medical purpose under Article 2(1) |
| 2 | Does ARISE provide information used to take decisions with therapeutic purposes? | Yes. The clinical dashboard provides KPIs and error detections used by the therapist to decide treatment progression. The real-time biofeedback directly modifies the patient's motor execution, which is itself a therapeutic intervention |
| 3 | Could such decisions cause death or irreversible deterioration of a person's state of health? | No. The therapist supervises every session, the device is contactless, and the worst foreseeable consequence of a misclassification is a sub-optimal repetition that the therapist corrects |
| 4 | Could such decisions cause serious deterioration or a surgical intervention? | No, for the same reasons as step 3 |
| 5 | Is ARISE intended for monitoring of vital physiological parameters where variations could result in immediate danger? | No. The kinematic variables observed (joint angles, velocities) are not vital signs |
| Result | Default Rule 11 outcome | Class IIa |

### 4.3 Consequences of the confirmed classification

The Class IIa confirmation has the following consequences across the project.

| Area | Consequence |
| --- | --- |
| Conformity-assessment route | Annex IX, requiring a Notified Body |
| Notified Body involvement | Required for both QMS audit (Annex IX Chapter I) and technical documentation assessment (Annex IX Chapter II) |
| Clinical evaluation | Full clinical evaluation per MDR Article 61 and Annex XIV Part A, including a Clinical Evaluation Plan, a Clinical Evaluation Report, and a Post-Market Clinical Follow-up (PMCF) plan |
| Quality management | Certified ISO 13485 quality management system before Notified Body audit |
| Technical documentation | Full Annex II content set, plus Annex III post-market surveillance plan |
| Vigilance | Article 87 reporting timelines apply. PSUR cadence per Article 86 is every two years for Class IIa |
| Phase 2 timeline | 12 to 18 months from end of Phase 1 to CE marking, dominated by Notified Body queue and audit cycle |

## 5. Conformity assessment pathway

### 5.1 Chosen route

The conformity-assessment route for the Class IIa ARISE device is MDR Annex IX, which combines:

Annex IX Chapter I: assessment of the quality management system (QMS) by the Notified Body

Annex IX Chapter II: assessment of the technical documentation by the Notified Body, on a representative sample basis for Class IIa devices

The alternative route, Annex XI Part A (product verification on the basis of type-examination), was considered and rejected. Annex IX is preferred because:

It provides broader QMS validation, which supports the Equipment-as-a-Service commercial model where the same QMS underpins multiple deployment configurations

It avoids the need for a separate Annex X type-examination, which Annex XI Part A would presuppose for this class of device

It is the route most commonly chosen by software-as-a-medical-device manufacturers for Class IIa devices, and Notified Bodies are most familiar with this pathway for software

### 5.2 Notified Body engagement plan

| Milestone | Target timing | Activity |
| --- | --- | --- |
| Notified Body shortlist | M14 (Phase 1) | Identify candidate Notified Bodies designated for MDR with the relevant codes for software in rehabilitation. Candidate codes include MDA 1207 (active devices for rehabilitation) and MDS 1010 (devices incorporating software). The final code matrix is confirmed in Phase 2 |
| Pre-submission meeting | M18 to M20 (Phase 1 end) | Informal meeting with the preferred Notified Body to align on the technical-documentation scope and on the QMS readiness expectations |
| Notified Body selection | First quarter of Phase 2 | Signed contract with the chosen Notified Body |
| QMS audit | Second quarter of Phase 2 | Annex IX Chapter I assessment |
| Technical documentation assessment | Third quarter of Phase 2 | Annex IX Chapter II assessment |
| Issue of EC certificate | Fourth quarter of Phase 2 to first quarter of the following year | Subject to the Notified Body queue, which is the dominant timing uncertainty |
| CE marking and commercial launch | Following EC certificate issue | Phase 2 EaaS launch |

# Part III: Engineering and clinical compliance

## 6. Risk management

The risk-management process follows ISO 14971:2019 and is conducted by the Innovina internal process for medical-device projects. The process is conducted to a depth appropriate to Class IIa under Annex IX and is subject to review by the Notified Body in Phase 2.

### 6.1 Risk-management process

| N | Activity | Output |
| --- | --- | --- |
| RM1 | Risk-management plan | Document defining the process, responsibilities, acceptability criteria |
| RM2 | Hazard analysis | Identification of foreseeable hazards across normal use and reasonably foreseeable misuse |
| RM3 | Risk evaluation | Per-hazard estimation of severity and probability, against acceptability criteria |
| RM4 | Risk control | Mitigation measures: inherent safety, protective measures, information for safety |
| RM5 | Residual risk evaluation | Post-mitigation assessment and sign-off |
| RM6 | Risk-management report | Summary submitted to the Notified Body as part of the Annex IX Chapter II technical documentation assessment |

### 6.2 Identified hazards (preliminary)

The full risk register is maintained as a separate controlled document, available to the Notified Body during the Annex IX audit. The hazards below are the high-level inventory.

| ID | Hazard | Mitigation category |
| --- | --- | --- |
| H1 | Patient fall during exercise | Information for safety (Instructions for Use mandate therapist supervision), always-on G1 safety flag in software |
| H2 | Misclassification of correct rep as incorrect, leading to confusing biofeedback | Protective measure: multi-layer detection (rule + per-patient z-score), confidence gating |
| H3 | Misclassification of incorrect rep as correct, missing a clinical concern | Protective measure: multi-layer detection reduces single-point failure |
| H4 | Therapist over-reliance on dashboard analytics | Information for safety: clear labelling that data is decision-support, not decision-making. Therapist training covers this |
| H5 | Device malfunction during session (camera failure, software crash) | Inherent safety: graceful degradation. Protective measure: session-resume policy and on-site spare units |
| H6 | Cybersecurity incident exposing patient data | Cross-reference to DPIA Section 6 and DMP Section 8. Compliant with MDR Annex I GSPR 17.2 |
| H7 | Misuse on patient outside intended population (severe cognitive impairment, acute instability) | Information for safety: contraindications explicit in IFU. Protective measure: therapist enrolment gating |
| H8 | Inaccurate AI inference due to environmental factors (lighting, clothing, occlusion) | Inherent safety: rigidification reduces noise. Protective measure: confidence threshold suppresses low-quality reps |

## 7. Clinical evaluation strategy

### 7.1 Phase 1 clinical investigation

The Phase 1 demonstration in WP5 is a clinical investigation under MDR Articles 62 to 82, conducted under an Ethics-Committee-approved protocol (D1.2). The protocol is the controlling document for how the investigation is run.

| Phase 1 clinical-evaluation component | Source |
| --- | --- |
| Clinical Investigation Plan | D1.2 |
| Investigator's Brochure (referencing intended purpose, risks, prior data) | D1.2 companion document |
| Recording of adverse events under Article 80 | DMP Section 3.3 (D-TRIAL AE sub-component), DPIA Section 6.9 |
| Performance against acceptance criteria | Validation Acceptance Criteria document, finalised in WP5 |
| Usability evaluation | DMP Section 3.3 (D-USABILITY), D5.1 |
| Safety evaluation | DPIA Section 5 risk register, plus the ISO 14971 risk-management report |
| Patient and therapist feedback | T5.3, D5.1 |

### 7.2 Phase 2 clinical evaluation

For a Class IIa device, full clinical evaluation per MDR Article 61 and Annex XIV Part A is required. The clinical-evaluation chain in Phase 2 comprises the documents below, prepared in accordance with MDCG 2020-5 (Clinical Evaluation, Equivalence) and MDCG 2020-6 (Sufficient Clinical Evidence for Legacy Devices) and the relevant MEDDEV 2.7/1 rev 4 guidance where still applicable.

| Phase 2 document | Purpose |
| --- | --- |
| Clinical Evaluation Plan (CEP) | Defines the scope of the clinical evaluation, the equivalent or comparable devices considered, the literature-review strategy, and the acceptance criteria for clinical safety and performance |
| Literature review | Systematic review of the published evidence on markerless biomechanical assessment, on biofeedback in motor rehabilitation, and on the Sit-to-Stand task, with appraisal of each source's relevance and quality |
| Clinical Evaluation Report (CER) | Compiled from D5.1 (the TRL6 Validation Report) plus the literature review, with the benefit-risk analysis required by MDR Article 61 |
| Post-Market Clinical Follow-up (PMCF) Plan | Per Annex XIV Part B. Mandatory for Class IIa devices. Defines the proactive collection and evaluation of clinical data from the device in routine clinical use post-CE marking |

The depth of the clinical evaluation matches the Class IIa expectation. Equivalence claims against existing devices are evaluated under MDCG 2020-5 and are unlikely to be sufficient on their own, given the novel combination of markerless RGB capture with biomechanical AI inference. The clinical evaluation therefore relies primarily on the data generated by the Phase 1 clinical investigation plus the planned PMCF.

## 8. Quality management

Innovina operates an internal quality management process during Phase 1 with the target of achieving certified ISO 13485:2016 compliance before the Notified Body QMS audit in Phase 2. The process covers seven areas.

| QMS area | Phase 1 alignment | Phase 2 readiness |
| --- | --- | --- |
| Document and change control | All MDR-relevant artefacts under controlled change history | Audit-ready for ISO 13485 clause 4.2 |
| Design and development controls | WP3 and WP4 workflows produce traceable design history | Audit-ready for ISO 13485 clause 7.3 and IEC 62304 |
| Procurement controls | Coach hardware procurement traceable | Audit-ready for ISO 13485 clause 7.4 |
| Customer-property handling | Patient data covered by DMP and DPIA | Audit-ready for ISO 13485 clause 7.5.4 |
| Internal audit | Six-month cycle during Phase 1 | Audit-ready for ISO 13485 clause 8.2.4 |
| Management review | Six-month cycle during Phase 1 | Audit-ready for ISO 13485 clause 5.6 |
| Corrective and preventive action (CAPA) | CAPA process operational from M6 | Audit-ready for ISO 13485 clauses 8.5.2 and 8.5.3 |

# Part IV: Lifecycle and reporting

## 9. Technical documentation (MDR Annex II and Annex III)

Phase 1 produces, in Annex II-compatible structure, the technical documentation listed below. Each item has controlled change history. The total document set is the full content set for Class IIa under Annex IX Chapter II.

| Annex II Section | Phase 1 status | Phase 2 use |
| --- | --- | --- |
| Section 1, Device description and specification | Drafted in D1.1 + D2.1B + D3.3 | Foundation of the Annex II technical file |
| Section 2, Information supplied with the device (IFU, labelling) | First draft at M18 | Finalised pre-CE submission |
| Section 3, Design and manufacturing information | Maintained in WP3 (D3.3) and WP4 (D4.1) | Annex II Section 3 |
| Section 4, General Safety and Performance Requirements (GSPR) checklist | Drafted M9 to M18 | Annex II Section 4 |
| Section 5, Benefit-risk analysis and risk management | Per Section 6 of this plan plus the ISO 14971 risk-management file | Annex II Section 5 |
| Section 6, Product verification and validation | D5.1 plus the bench-test reports and the software verification per IEC 62304 | Annex II Section 6 |
| Section 6.1(e), Clinical evaluation | D5.1 plus the Phase 2 Clinical Evaluation Report | Annex II Section 6.1(e) |
| Annex III, Post-Market Surveillance Plan | Drafted M22, finalised M24 | Submitted as part of the Annex IX Chapter II technical-documentation assessment |

## 10. Vigilance and post-market surveillance

### 10.1 Phase 1 obligations (clinical investigation)

| Obligation | MDR reference | Implementation |
| --- | --- | --- |
| Record adverse events | Article 80(1) | D-TRIAL AE sub-component (DMP Section 3.3, Section 4.1), AE schema in DMP Section 5.4 |
| Report serious adverse events to competent authority | Article 80(2) | Procedure defined in D1.2 Section 12. Reporting timeline: immediately for events that present an imminent risk of death, serious injury, or serious illness; otherwise within 7 days |
| Inform Ethics Committee of safety-significant events | Article 80(3) | Procedure defined in D1.2 Section 12 |
| Implement halt criteria | Article 78 | Halt criteria defined in D1.2 Section 14 |
| Maintain trial master file | Article 72(3), Annex XV Chapter III, point 3 | Innovina retains. Retention period at least 10 years from the end of the clinical investigation, available for inspection |

### 10.2 Phase 2 obligations (post-market, Class IIa)

Phase 1 does not have a commercial product on the market, so there is no formal post-market surveillance obligation. However, the architecture for PMS is established in Phase 1 so that it activates immediately on CE marking. The reporting cadence for Class IIa is set out below.

| Phase 2 obligation | MDR reference | Class IIa specifics | Phase 1 preparatory work |
| --- | --- | --- | --- |
| Post-market surveillance plan | Article 84, Annex III | Plan in place at CE marking | Architecture established in DMP Section 3.5 (continuous activities) |
| Periodic Safety Update Report (PSUR) | Article 86 | Updated at least every two years for Class IIa, made available to the Notified Body through EUDAMED | Reporting templates drafted at M22 |
| Vigilance: reporting of serious incidents and Field Safety Corrective Actions (FSCAs) | Article 87 | Serious incidents reported without delay and not later than 15 days; for incidents presenting a serious threat to public health, immediately and not later than 2 days; for death or unanticipated serious deterioration, not later than 10 days | Procedure drafted at M22 |
| Trend reporting | Article 88 | Statistically significant increases in non-serious incidents or expected undesirable side-effects must be reported | KPI-distribution drift monitoring (DMP Section 7.3) repurposed for trend detection |
| PMCF plan and reports | Annex XIV Part B | Mandatory for Class IIa. PMCF report at least every two years and integrated into the PSUR | Drafted in Phase 2 |
| EUDAMED registration | Articles 27, 28, 29, 31 | UDI assignment, manufacturer registration, device registration before placing on the market | UDI structure prepared in Phase 1 |

## 11. Phase 1 to Phase 2 transition

The transition is gated by completion of D5.1 at M24. Phase 2 begins with the steps below.

| Phase 2 milestone | Trigger | Class IIa timing |
| --- | --- | --- |
| Confirm classification (already done in Phase 1) | n/a | Confirmed in this document, Section 4 |
| Appoint PRRC and confirm QMS | D5.1 sign-off | Q1 of Phase 2 |
| Sign Notified Body contract | QMS in place, pre-submission complete | Q1 of Phase 2 |
| Compile Annex II technical file | Conformity assessment procedure | Q1 to Q2 of Phase 2 |
| Notified Body QMS audit (Annex IX Chapter I) | Contract and QMS in place | Q2 of Phase 2 |
| Notified Body technical-documentation assessment (Annex IX Chapter II) | Technical file submitted | Q3 of Phase 2 |
| EC certificate issued | NB assessments complete | Q4 of Phase 2 or Q1 following year |
| CE marking and EaaS commercial launch | EC certificate in hand | Following EC certificate |
| First PSUR | 24 months after first device on the market | Per Article 86 cadence |
| First PMCF report | 24 months after first device on the market | Per Annex XIV Part B |

# Part V: Governance

## 12. Roles and responsibilities

| Role | Organisation | MDR responsibility |
| --- | --- | --- |
| Project Lead | Innovina | Overall MDR compliance. Sign-off on this plan |
| Person Responsible for Regulatory Compliance (PRRC, Article 15) | Innovina | Overall MDR responsibility per Article 15. Sign-off on conformity assessment submissions |
| Quality Manager | Innovina | QMS alignment with ISO 13485, document control, CAPA |
| Risk Manager | Innovina | Risk-management process owner per ISO 14971. Maintains the risk register |
| Clinical Lead | Studio Buccarella | Conduct of the clinical investigation. AE recording per Article 80 |
| Scientific Lead | DINOGMI | Independent scientific validation of clinical-performance claims. Gold-standard reference instrumentation at REHELAB |
| Notified Body | To be selected in Phase 2 | Annex IX Chapter I QMS assessment and Annex IX Chapter II technical-documentation assessment for the Class IIa device |
| Italian National Competent Authority | Ministero della Salute, Direzione Generale dei Dispositivi Medici e del Servizio Farmaceutico | Notification of the Phase 1 clinical investigation under Article 70. Vigilance reporting destination in Phase 2 |